PulseGuard  — a multi-event intraoperative early-warning system

Technical summary  ·  Aastik Rajan, Johns Hopkins University  ·  aastikc15@gmail.com

Live demo: aastikrajan.github.io/pulseguard    Code + full benchmark: github.com/AastikRajan/pulseguard

Research prototype — not clinical software. Every figure below is out-of-sample, on a single locked test split.

PulseGuard watches all of a patient's vital signs during surgery and warns, minutes ahead, that they are about to deteriorate — saying which of nine problems is coming and why. It was built solo on public data (VitalDB, 5,151 surgeries from Seoul National University Hospital). Most of the effort went not into the model itself but into whether it can be trusted: calibrating its alarms, testing whether it survives a move to a different hospital, mapping exactly where it fails, and separating real causal effects from confounding.

Data and evaluation

- VitalDB: 5,151 surgeries, 35 signals at 1 Hz (plus 500 Hz waveforms).

- Locked split (seed fixed once): 20% of cases held out = 1,030 test cases / 303,909 time windows. No patient appears in both training and test.

- Labels come from the recorded vitals and events; nothing below is tuned on the test set.

What the system does

| Capability | Result (held-out test set) |
| --- | --- |
| Unified early-warning — any deterioration, 5-min horizon | AUPRC 0.454 · AUROC 0.849 · ~4.3-min median lead (base rate 12%) |
| Alarm precision after conformal calibration | 62–71% real vs ~12% for deployed threshold monitors; ~300× fewer alarms |
| Multi-event detection — 9 distinct events | per-event AUROC 0.74–0.96 |
| Cross-hospital, zero-shot (no retraining) | AUROC 0.78–0.80 Seoul → UC-Irvine |
| Causal: does low blood pressure cause kidney injury? | RR 1.77 → 1.07 mostly confounding |

The nine events: hypoxemia, hypotension, bradycardia, tachycardia, bronchospasm, hypercapnia, ST-deviation (ischemia), light anesthesia (awareness), and apnea. Each is detected by its own model — a single 9-way classifier does worse, because these events often happen together.

How it works, in brief

- Every 30-second window of vitals is turned into features — current levels, short-term trends, variability, cross-signal coupling, and leak-free event-history (what has already happened in this case). A gradient-boosted model scores the chance of deterioration in the next five minutes.

- Raw scores are turned into trustworthy alarms by conformal calibration (which gives a precision target), then episode-grouping collapses a burst of firings into a single alert — that's the ~300× reduction.

- For cross-hospital use the model keeps only the channels every monitor has (heart rate, SpO2); adding rich, source-specific channels helps within one hospital but hurts transfer to another.

Causal sub-study — does hypotension cause kidney injury?

A separate arm asks whether intraoperative low blood pressure causes postoperative acute kidney injury. The crude association is RR 1.77. After deconfounding for recorded patient factors (g-formula), it falls to RR 1.07 — about 90% of the apparent effect is confounding by indication (sicker patients get both). The null replicates on a second hospital and matches the 2025 randomized trials, where preventing hypotension did not prevent kidney injury.

Where it can't be trusted (the boundary I mapped)

- Intraoperative hypotension is close to unpredictable — system recall ~0.12. I confirmed this seven independent ways (operating-point analysis, lead-signal tests, critical-slowing-down, a dynamical network biomarker, short-horizon reachability, phenotype analysis, and 500 Hz heart-rate-variability), so I report it as a hard limit instead of tuning around it.

- Four signals carry almost all the information: heart rate, SpO2, the pressure line, and the ventilator. Heart rate and SpO2 have no substitute — so a low-cost version using only ECG + pulse-oximetry is plausible where arterial lines aren't available.

Why the numbers are believable

- One locked test split, chosen once and never tuned on.

- I caught a data leak in my own pipeline that inflated the headline (0.48 → 0.454 once fixed) and quarantined it.

- Every null and failure is reported, not buried — the honest boundary above is part of the result.

What's new, and what I'm looking for

I haven't found a published system that combines all of multi-event early-warning, zero-shot cross-hospital transfer, a causal analysis, and calibrated alarms — alongside an explicit map of what it cannot predict. The contribution is integration, interpretability, and rigor rather than a new biomarker. I'm looking for feedback, a research role, or a collaborator to take the strongest pieces — the cross-hospital transfer and the honest benchmark — toward a paper and external validation.